Suitability of
ALPAO deformable mirrors
for open-loop control
Urban Bitenc
Durham University
Workshop on Real-Time Control for Adaptice Optics, 4th edition
2016 Dec 21

---
Contents
- ALPAO DMs tested
- Two issues found:
  - creep
  - thermal effect
- Software compensation developed for both
- Conclusions

---
Open-loop control
OPEN LOOP
CLOSED LOOP
DM
DM
Figures stolen from
Chao Li et al. 2009,
and modified to fit this talk.
e.g. Multi-Object AO

---
When is a DM good for open loop?
When
the surface shape
responds to
the actuator commands
exactly in the way you expect.
- Response is linear
- Response is unique:
      - same commands --> same DM shape
- Method: DM in front of an interferometer
STUDIED FOR THIS TALK

---
Monitor DM surface over night
Flatten the DM and observe how its shape evolves with time.
Measure the DM shape every 6 minutes and plot its RMS.
The shape drifts away from the initial shape and stabilizes only after several hours.

---
DM shape for 6
hours
DM 1 hour
after flattening
DM flattened
DM many hours
later
DM COMMANDS A
DM COMMANDS B
BEFORE
AFTER
- The DM surface remembers its shape and drifts after a change.
- This is a long-term effect:
  - flattening takes 15 minutes and afterwards the effect is still there
  - takes hours to reach the equilibrium shape
        CHANGE FROM A TO B

---
Compensate for creep
IDEA:
 =
  + x(t)*
Shape B(t)
Shape A
Shape B(0)
(Shape before the change)
(Shape immediately after the change)
(Shape some time after the change)
(1) Measure the correction facotrs:
Correction factors
Interferometer images
Several iterations,
12 hours each
(2) Apply the correction:

---
Applying creep compensation
NO COMPENSATION
WITH COMPENSATION
ZOOM IN
ZOOM IN
RMS change in 3 hours:
80 - 90 nm (temperature dependent)
RMS change in 3 hours:
2.5 - 6 nm (temperature dependent)
Optics Express Vol. 22, Iss. 10, pp. 12438–12451 (2014)

---
For a more general use of the DM
- Change DM shape several times:
- The DM will “remember” all these shapes
- Correct for all shapes from the past few hours:
B0
t
t0
t1
t2
B1
B2
B3
tB0
tB1
tB2
DB1,0
DB2,1
DB3,2

---
Performance
Compensating after 7 shape changes:
NO COMPENSATION: WITH COMPENSATION:
        90 - 110 nm RMS 6 - 10 nm RMS

---
Performance
Compensating the 6 intermediate shapes:
NO COMPENSATION: WITH COMPENSATION:
        90 - 200 nm RMS 7 - 13 nm RMS

---
- Creep compensation performs very well
- Result: stable DM surface

---
THERMAL EFFECT
- The DM shape changes if the DM warms up
- Currents running through actuator coils warm up the DM

---
Amplitude vs power
The amplitude of the thermal effect depends on the power the actuators consumed in the last ~2 minutes
actuator commands,
between [-1.0, 1.0]
N = 241, 277 (number of actuaotrs)

---
Software compensation

---
Conclusion
- The ALPAO DM exhibits creep behaviour
  - It takes hours to stabilize.
- ALPAO DM changes shape if it warms up
- One can successfully compensate for this in software
- [1] Optics Express Vol. 22, Iss. 10, pp. 12438–12451 (2014)
- [2] Submitted to Optics Express
NO COMPENSATION:
        90 - 110 nm RMS
WITH COMPENSATION:
        6 - 10 nm RMS

---
BACKUP SLIDES

---
How stable and repeatable is the DM shape
ACTUATOR COMMANDS
THE CORRESPONDING DM SHAPE

---
A few hours later...
DM SHAPES FOR ACTUATOR COMMANDS
FROM THE PREVIOUS SLIDE
scary